Section 340.540  Calibration of Survey Instruments

a)	Unless specified in another Part, a licensee shall have each survey instrument used to show compliance with this Part calibrated before first use, annually and following a repair that affects the calibration.  A licensee shall:

1)	Calibrate all scales with readings up to 10 mSv (1000 mrem) per hour with a radiation source;

2)	Calibrate two separated readings on each scale or decade that will be used to show compliance; and

3)	Conspicuously note on the instrument the date of calibration.

b)	A licensee may not use survey instruments if the difference between the indicated exposure rate and the calculated exposure rate is more than 20 percent.

(Source:  Added at 29 Ill. Reg. 20841, effective December 16, 2005)