Section 350.APPENDIX B  General Requirements for Inspection of Industrial Radiographic Equipment

a)	Panoramic devices (devices in which the sealed source is physically removed from the shielded container during exposure) shall be inspected for:

1)	Radiographic Exposure Unit

A)	Abnormal surface radiation levels anywhere on camera, collimator or guide tube;

B)	Condition of safety plugs;

C)	Proper operation of locking mechanism;

D)	Condition of pigtail connector;

E)	Condition of carrying device (e.g., straps, handle, etc.); and

F)	Proper labeling.

2)	Source Guide Tube

A)	Rust, dirt or sludge buildup inside the source tube;

B)	Condition of source tube connector;

C)	Condition of source stop;

D)	Kinks or damage that could prevent proper operation; and

E)	Presence of radioactive contamination.

3)	Control Cables and Drive Mechanism

A)	Proper drive mechanism with camera, as appropriate;

B)	Changes in general operating characteristics;

C)	Conditions of connector on drive cable;

D)	Drive cable flexibility, wear and rust;

E)	Excessive wear or damage to crank assembly parts;

F)	Damage to drive cable conduit that could prevent the cable from moving freely;

G)	Connection of the control cable connector with the pigtail connector for proper mating;

H)	Proper operation of source position indicator, if applicable; and

I)	Presence of radioactive contamination.

b)	Directional beam devices containing radioactive material shall be inspected for:

1)	Abnormal surface radiation;

2)	Changes in the general operating characteristics of the device;

3)	Proper operation of shutter mechanism;

4)	Chafing or binding of shutter mechanism;

5)	Damage to the device which might impair its operation;

6)	Proper operation of locking mechanism;

7)	Proper drive mechanism with camera, as appropriate;

8)	Condition of carrying device (e.g., strap, handle, etc.); and

9)	Proper labeling.

c)	X-ray equipment shall be inspected for:

1)	Change in the general operating characteristics of the device;

2)	Wear of electrical cables and connectors;

3)	Proper labeling of console;

4)	Proper console with machine, as appropriate;

5)	Proper operation of locking mechanism;

6)	Timer run-down cutoff;

7)	Damage to tube head housing that might result in excessive radiation levels; and

8)	Positive indication of x-ray production.

Source:   Added at 18 Ill. Reg. 7263, effective May 2, 1994)